Lawson Labs, Inc.  USB Data Acquisition Systems

LabView Sample application

Last update: 4-7-03

Win/XP Version

===========

OVERVIEW:

===========

Proper installation of hardware is required before use of sample application. If you haven't

yet done so you must first install the device driver for the hardware. If you received this file

from our website, then the INSTALL.ZIP file is the archive that contains everything required

for installation. The entire contents of that file should be placed onto a floppy diskette to be

used for installation. Please read the Install.DOC file located in that archive for complete

installation information.

Sample application currently available only in development environment, there is no compiled

executable LabView type application currently available. Although it has some  "roughing in"

for use of multiple devices it currently only fully supports one device at at time. Our DLL and

DevceDriver has support for 32 of our USB devices connected and running at the same time.

The sample application is very well documented so please be sure to read all segments of

documentation on main panel as well as very extensive documentation within each of the

sequences within the actual  "Diagram"  source code itself. It can become very complicated

very quickly if each label of documentation isn't read as you follow through the application

source code.

This application requires the following files to run:

in  /Windows/system32/Drivers  directory:

LL_USB.SYS

in  /Windows/  directory:

LL_USB.INI

in application directory:

LL_USB2k.DLL	- The DLL

LL_USB.vi	- Main application panel

DevStat.vi	- Handling of device status information retrieval

Global.vi		- All Global variables

Read.vi		- Used to access DLL for direct reading of device

Write.vi		- Used to access DLL for direct writing of device

DLL Functions used by sample application:

BOOLEAN  SP_DisableDialog(void);

Prevents DLL from displaying a dialog box used for runtime messaging

as well as handling of USB device connect/disconnect system messages.

LabView and some other application interfaces have problems with the

Dialog's resources occupying their program space and will crash on exit

if the Dialog box is used. ALWAYS call this function FIRST THING when

the program first starts.

BOOLEAN   SP_InitAllDevices(

USHORT* pDevList, BYTE* pbNumDevices_LV,

DOUBLE dblRate, BOOLEAN fShowDevList);

This call is used to initialize all devices. Unlike typical serial devices, USB

type devices require an enumeration process. After that, each device needs

to be initialized to set it to certain configuration parameters. Place all of

the device IDs that you wish to initialize in the pDevList array. The number

of devices is limited to 32. The DLL will attempt to initialize whichever IDs

are in that list. Any devices that fail to initialize will be removed from the list.

On return, the list will contain only valid device ID's that initialized properly.

If scanning, do not place the devices that are currently scanning in the list

since the DLL will automatically attempt to close and reenumerate any

device that is currently in the list including the ones that are scanning if they

are in the list.

BOOLEAN  SP_GetOneConversion(

USHORT usDevID, UINT* uiConvRawCount,

BYTE bCurChan, BYTE* pbLastDigin);

The DLL will get read the data for the selected channel from the board.

It returns the data in a  "Raw Count"  24-bit data format, so it is up to the

application to convert that to a voltage. A suggested technique for doing

that is demonstrated in the sample application code. The DLL will also send

back the current digital input that it reads from the board. The DLL retrieves

the digital input value automatically anytime the board is read.

BOOLEAN  SP_SendDigOut(USHORT usDevID,

BYTE bDigOut, BYTE* pbLastDigin);

The DLL will send the digital output value to the board and return the last

digital input value that was read.

BOOLEAN  SP_StartScan(

USHORT usDevID, DOUBLE* dblScanObject);

The DLL will start the board(s) scanning. Once started, the data can be

read using the following function call. There are three scan modes which

are discussed in more detail at the bottom of this page.

BOOLEAN  SP_ReadScanData(

USHORT usDevID, USHORT numPointsToRead,

UINT* DynaVoltageBuff, DOUBLE* dblScanObject);

The DLL will read the scan data and return the data within the array that is

passed in the call. Although the array name indicates that a voltage is to be

returned, the data returned is actually returned as particular packets that

consist of data points as well as sometimes digital input data as described

at the end of this page.

BOOLEAN   SP_ReadScanDataWithDigin(

USHORT usDevID, UINT* pDynaVoltageBuff,

BYTE* pbDiginBuff, BYTE* pbCurChan,

USHORT numPointsToRead, DOUBLE* dblScanObject);

The DLL will read the scan voltage data as well as the digital input value

that was read the same time as the particular voltage. This scan type is

particular useful if you do not want to start recording the voltage data until

a particular trigger value has been detected by way of the digital input value.

This call requires multiple buffers to be passed in the call and although it isn't

currently demonstrated in the sample application source code, it is very well

documented at the bottom of this page.

BOOLEAN  SP_EndScan(USHORT usDevID);

The DLL will end the scanning. Call this function when you want the scan

process to end.

BOOLEAN  SP_GetDevStat(

USHORT usDevID, UINT* _uiDevStat);

The DLL will return the current status of one device. Occasionally I/O errors

can occur. For example the USB cable might be disconnected from the

device. This call will return such information to the caller.

BOOLEAN  SP_GetDevStatEx(UINT* _uiDevStat,

USHORT* _pusAllDevList, BYTE* bNumDevices);

The DLL will return the current status of multiple devices. Occasionally I/O

errors	can occur. For example the USB cable might be disconnected from

the device. This call will return such information to the caller.

BOOLEAN  SP_SendRate(

DOUBLE* dblRate, USHORT usDevID);

The DLL will set the device to run at a particular rate.

BOOLEAN  SP_SystemCalibration(

USHORT usDevID, BYTE bCurChan);

The DLL will do a system full scale and offset calibration of the device.

BOOLEAN  SP_SendDAC(USHORT usDevID,

DOUBLE dblVoltage,  BYTE bDAC, BYTE* bDigIn);

The DLL will send the passed analog output to the device. This is only

available for Model-302 devices.

BOOLEAN  SP_SendChan( USHORT usDevID,

BYTE bCurChan,  BYTE* pbLastDigin);

The DLL will set the device channel to the channel that is passed.

===========

SCANNING:

===========

The Model 301 and 302 have 6 scanning modes. The sample application

shows how to use 2 of them. The available modes as listed within the DLL's

header file are:

*single chan regular		-  0x89

single chan with digin 		-  0x8D

*multi chan regular		-  0x8B

multi chan with cal			-  0x8C

multi chan with digin		-  0x8E

multi chan with cal and digin	 -  0x8F

*used in sample application

Please see sample application for use of regular Single-chan as well as

the Multi-chan scan modes for an overview of the scanning process. What

follows is a better explanation of the parameters used as well as how the

data will be returned by the call to  SP_ReadScanData(...).

The DLL uses structured data for interraction between the DLL, application,

and device driver during the scan process. For the LabView interface, the

structured data is passed as a 32-member array of doubles. The doubles

are then converted to other data types by the DLL before usage. Below is

how the structured data is arranged by the DLL as well as the definition of

the current read status as returned by the DLL with each call by the app to

read the data.

The DLL uses the followig structured data format:

typedef struct _SCAN_OBJECT_DIO {

BYTE 		SO_bScanType;			// (APP WRITES) Definitions below

#define	SINGLE_CHAN_SCAN				1

#define	MULTI_CHAN_SCAN				2

#define	MULTI_CHAN_CAL_SCAN			3

#define	SINGLE_CHAN_DIGIN_SCAN			4

#define	MULTI_CHAN_DIGIN_SCAN			5

#define	MULTI_CHAN_CAL_DIGIN_SCAN		6

DOUBLE	SO_dblRate;			// (APP WRITES) device rate

UINT		SO_uiTotalPointsReadByDrvr;	// (DLL WRITES) points read by driver.

UINT		SO_uiPointsInBuffer;		// (DLL WRITES) points in DLL/Driver

shared buffer.

UINT		SO_uiStatus;			// (DLL WRITES) Definitions below

#define SCAN_STATUS_UKNOWN		0x00	// Driver status unknown

#define SCAN_RUNNING 			0x01	// Driver scan thread is running

#define SCAN_DATA_REQUEST_SUCCEEDED	0x02	// Pending driver to device

data request succeeded

#define SCAN_ENDING 				0x04	// DLL has requested driver to

end scanning

#define SCAN_HALTED 				0x08	// Driver scan thread was

halted (possible I/O err,

device unplug, etc)

#define SCAN_CHECKSUM_ERROR		0x10	// Driver's data checksum

doesn't agree with device's

checksum

#define SCAN_VOLTAGE_BUFFER_WRAP	0x20	// The circular data buffer shared

between DLL and Driver has

wrapped

#define SCAN_MICRO_CODE_BUFF_WRAP	0x40	// The device's (hardware)

internal buffer has wrapped

#define SCAN_DEVICE_IO_ERROR		0x80	// Driver received an I/O error

#define SCAN_DEVICE_REMOVED		0x100	// The USB cable has been

disconnected

#define SCAN_DEVICE_NOT_FOUND		0x200	// The DLL was unable to find

the requested device to

begin the scan

UINT		SO_uiSizeVoltageArray;		// (APP WRITES) The size of app's voltage array.

UINT		SO_uiCurHead;			// (DLL WRITES) The device driver increments

whenever it adds a byte to DLL/Driver buffer

UINT		SO_uiCurTail;			// (DLL WRITES) The DLL increments this

whenever it reads a byte DLL/Driver buffer

BYTE		SO_bDigitalInput;		// (DLL WRITES) Current digital input value. Driver

updates this every 10th point read

BYTE		SO_bChecksum;		// (DLL WRITES)  Driver's current checksum

BYTE		SO_bDeviceDisconnect; 	// (DLL WRITES) required but value is ignored

(backward compatibility)

clock_t	SO_clkScanStartTime;		// not currently changed

clock_t	SO_clkScanEndTime;		// not currently changed

BYTE		SO_bScanArg;			// (APP WRITES) Number of channels to scan.

} SCAN_OBJECT_DIO, *PSCAN_OBJECT_DIO;

The DLL Reads the requested number of points (3 bytes per point)  from the device. If

the requested number of points is unavailable, the DLL returns FALSE. Otherwise, the

DLL returns TRUE. All  SO_uiStatus  Flags are  OR'd on by the DLL. It is best to call this

function as frequently as possible, and inspect the flags with each call in case there was

an I/O error in the device. The DLL will also return FALSE if an error occurred during the

read.

Note:  Application's dblScanObjectArray  array members that represent the

SCAN_OBJECT_DIO structure members above are only updated when this function is

called.

If upon return from this call, you discover that  SO_uiPointsInBuffer is growing too quickly,

it is then most efficient to quickly make multiple calls to this function with requests for a

larger number of points.

=================

Note:

=================

There are 4 entries in  LL_USB.INI  to help improve driver capability to maintain stability

against other drivers that may behave incorrectly. Each of the entries are explained

in detail in the "iniFile.htm"  file found in the API documentation provided in HTML

format which can be read directly from our website location:

http://www.lawsonlabs.com/Download/Manuals-Drivers/Model_301_302/Docs/API_Usage/index.htm

Downloaded as an archive:

http://www.lawsonlabs.com/Download/Manuals-Drivers/Model_301_302/Docs/API_Usage/archive/API_USB.zip

or just the iniFile.htm can be viewed here:

http://www.lawsonlabs.com/Download/Manuals-Drivers/Model_301_302/Docs/API_Usage/IniFile.htm

Handling of buffered data:

There are special buffer conciderations when using the digital input type scans, since the

digital input that was recorded at the time a particular voltage was read is mixed in with

the data in the buffer. This type of scan is most useful for matching a particular trigger to a

voltage. For example you may not be interested in reading any voltage until a particular

digital input bit has been set.

The DLL returns the requested number of data points if available as well as the current

digital input with each point. See the information further up this page for explanation of the

structured data passed in the dblScanObjectArray array. The normal SingleChannel scan

is the most straight forward since it returns however many points that are requested in the

buffer that is passed. All other scans require special handling of the buffered data as

explained below. Please be sure to always request a number of points that represents

the number of channels that are being scanned. The device doesn't allow skipping of

channels so if you want to scan channel 0 and channel 3 for example, data will be returned

for 0,1,2,and 3. It is up to you to discard the data for the extra channels. Under the same

circumstance (4-channels), be sure to request a minimum of 4 points (or multiple thereof)

with each request to read data, except for the circumstance when you request 0 points

just to get the current status and number of points available as is explained within the

sample application source code. The rate at which the data is returned is also dependant

on the number of channels scanned and there is a settling time between each channel

change which requires us to throw away 5 points. So if you are scanning 2 channels at

1000Hz, the data would be returned at a rate of 100Hz. The Model-301 allows scanning

of 4 channels only while the Model-302 has 8 scan channels.

When using the Digital-Input scan option, you must also use a special DLL function call

SP_ReadScanDataWithDigin(...)   to retrieve that data. The function call requires and

additional buffer for the digital input values. The digital input array MUST be the same

dimension as the voltage data array. Always call this function with two array pointers. The

first is to receive the current 24-bit conversion, and the second is to receive the current

digital input. This is the function to call when you've used one of the Digital-Input scan

tokens to start the scan. The device will include a digital input with each data point. The

DLL will return the digital input that corresponds to the data point using the same index.

For example if you do a  MULTI_CHAN_DIGIN_SCAN  type scan, then DynaVoltageBuff[0]

will contain datapoint for channel 0 and pbDiginBuff[0] will contain the digital input that was

read at the same time that point was read. To further explain:

if you're doing a non-cal scan on 3 channels and you request 12 points and 12 points are

available, then on return your arrays will look like this:

DynaVoltageBuff [chan0, chan1, chan2, chan0, chan1, chan2, chan0, chan1, chan2,

chan0, chan1, chan2]

pbDiginBuff [diginForChan0, diginForChan1, diginForChan2, diginForChan0,

diginForChan1, diginForChan2, diginForChan0, diginForChan1,

diginForChan2, diginForChan0, diginForChan1, diginForChan2]

if you're doing a calibration scan on 3 channels and you request 12 points and 12 points

are available, then on return your arrays will look like this, where "offs" means channel 7:

DynaVoltageBuff [chan0, chan1, chan2, offs, chan0, chan1, chan2, offs,

chan0, chan1, chan2, offs]

pbDiginBuff [diginForChan0, diginForChan1, diginForChan2,  diginForChanoffs,

diginForChan0, diginForChan1, diginForChan2, diginForChanoffs,

diginForChan0, diginForChan1, diginForChan2, diginForChanoffs]

I've also added pbCurChan  to help in case you should get confused which channel you're

requesting data for. You can call this function (ST_ReadScanDataWithDigin()) anytime -

even with  numPointsToRead = 0 - to get that information. Your variable which that pointer

points to should always equal zero when you return from this call. If it does not, then the

number of points you're requesting isn't correct for the number of channels you've requested

to scan. That value is set on entry into the call, not on exit. Each read should start with

channel zero regardless of the number of channels being read.

===============

VERSION INFO:

===============

This date: 4-7-03

App:			4-7-03

DLL:			7,9,2003

DeviceDriver : 	7,9,2003

Hardware:		04